JURNAL SKALA HUSADA: THE JOURNAL OF HEALTH

Available online at: https://ejournal.poltekkes-denpasar.ac.id/index.php/JSH

Vol. x No. x Bulan 202x

P-ISSN 1693-931X I e-ISSN 2580-3700 Halaman: x-xx

Verification of Total Plate Count Testing Method on Dairy Goat Milk Referring to SNI 2897: 2008

Putu Ayu Suryaningsih1, Ida Bagus Oka Suyasa1*

1 Poltekkes Kemenkes Denpasar

Posted : Marc 30th,2025 ;  Reviewed : Apr 2nd,2025 ; Received: xxxxxxxx

ABSTRACT

Background: A method must be verified and validated because each laboratory has varying conditions. This diversity can occur due to differences in facilities, personnel competence, equipment and chemicals used. In addition, accredited testing laboratories have an obligation to carry out method validation and verification processes before implementing the method. Objective: In this study, the ALT testing method for dairy goat milk was verified using accuracy and precision test parameters from samples with low, medium, high and uncontaminated concentrations. Method: This research uses descriptive research, namely describing the results of verification of the ALT testing method on fresh goat's milk which refers to SNI 2897:2008. Results: The SNI 2897:2008 ALT testing method on dairy goat milk for all microbial concentrations can be applied at the Denpasar Ministry of Health Health Polytechnic Integrated Laboratory because the RSD and Percent Recovery values ​​do not exceed the limits set in SNI 2897:2008. Conclusion: The Total Plate Number testing method according to SNI 2897:2008 with a dairy goat's milk matrix can be applied as a testing method at the integrated laboratory of the Health Polytechnic of the Ministry of Health, Denpasar.

Keywords: verificationi, TPC, milk, SNI

ABSTRAK

Latar Belakang : Suatu metode harus diverifikasi dan divalidasi karena setiap laboratorium memiliki kondisi yang beragam. Keragaman tersebut dapat terjadi karena perbedaan sarana, kompetensi personel, peralatan, dan bahan kimia yang digunakan. Selain itu, laboratorium pengujian yang telah terakreditasi memiliki kewajiban untuk melakukan proses validasi dan verifikasi metode sebelum implementasi metode. Tujuan : Pada penelitian ini, dilakukan verifikasi metode pengujian ALT pada susu kambing perah dengan parameter uji akurasi dan presisi dari sampel dengan konsentrasi rendah, sedang, tinggi dan tidak terkontaminasi. Metode : Penelitian ini menggunakan penelitian deskriptif, yaitu menggambarkan hasil verifikasi metode pengujian ALT pada susu kambing segar yang merujuk pada SNI 2897:2008. Hasil : Metode pengujian ALT SNI 2897:2008 pada susu kambing perah semua konsentrasi mikroba  dapat diterapkan di Laboratorium Terpadu Poltekkes Kemenkes Denpasar karena nilai RSD dan Persen Recovery tidak melebihi batas yang ditetapkan dalam SNI 2897:2008. Kesimpulan : Metode pengujian Angka Lempeng Total sesuai SNI 2897:2008 dengan matriks susu kambing perah dapat diterapkan sebagai metode pengujian di Laboratorium terpadu Poltekkes Kemenkes Denpasar.

Kata kunci: verifikasi, TPC, susu, SNI

- Corresponding Author:

Corresponding author's name

Affiliation and Institution

E-mail:

Copyright © 202x, Jurnal Skala Husada: The Journal of Health

INTRODUCTION

Standardization of sample testing is crucial for laboratories serving the general public. Laboratories play a vital role in research, education, and community service activities. Recognition of a laboratory's capability in testing a particular parameter can be achieved through a series of verification activities for a testing method. Testing laboratories should conduct verification activities for each method used in serving the public.

A method must be verified and validated because each laboratory operates under different conditions. These variations may arise due to differences in facilities, personnel competencies, equipment, and chemicals used. Generally, validation and verification of methods are carried out before a testing method is used in routine analysis within the laboratory. Furthermore, accredited testing laboratories are required to perform method validation/verification before method implementation (Clause 7.2 ISO/IEC 17025:2017).

Currently, the consumption of fresh goat milk is gaining popularity among the public. However, most dairy goat farms in Indonesia are still managed on a small scale using simple methods. Fresh goat milk can be a source of transmission for various pathogenic bacteria such as Staphylococcus aureus, Listeria monocytogenes, Campylobacter spp., and coliform bacteria. To date, fresh goat milk has received little attention from the government or other relevant institutions concerning quality control, safety, and hygiene (Nurliyani, 2015; Taufik, 2011).

The verified Total Plate Count (TPC) testing method at the Integrated Laboratory of Poltekkes Denpasar has been used to determine microbial counts in boiled eggs. The TPC testing method refers to SNI 2897:2008, which is the testing method for microbial contamination in meat, eggs, milk, and their processed products. Currently, no verified analytical method is available for microbial count testing in dairy goat milk. This study aims to verify the TPC testing method for dairy goat milk by evaluating accuracy and precision parameters for samples with low, medium, high, and uncontaminated concentrations. The verification results will be implemented as one of the testing services at the Integrated Laboratory of Poltekkes Kemenkes Denpasar. This effort serves as a part of public service to support the critical role of the Tri Dharma of Higher Education.

.

MATERIALS AND METHODS

This study employs a descriptive research method to describe the verification results of the TPC testing method on fresh goat milk, referring to SNI 2897:2008. The goat milk samples were obtained from a goat farm located at Jl. Maruti No.5, Pemecutan Kaja, North Denpasar District, Denpasar City, and sample testing was conducted at the Integrated Laboratory of Poltekkes Kemenkes Denpasar.

The equipment used includes an analytical balance (Radwag), a blender (Myako), a graduated pipette (Iwaki Pyrex®) 25 ml, a micropipette (Socorex) 100µl-1000µl and tips, a ball pipette (D&N), a measuring cylinder (Iwaki-Pyrex) 250 ml, sterile Petri dishes, test tubes (Iwaki), a biosafety cabinet (Biobase), an incubator (Esco), an autoclave (Tomy SX-500), a glass funnel, a glass stirrer, a magnetic stirrer, a colony counter, a sterile inoculation loop, a McFarland densitometer, a vortex mixer, a water bath, and a hot plate magnetic stirrer. The materials used include fresh goat milk, Plate Count Agar (PCA), and Buffered Peptone Water (BPW).

Each treatment group (uncontaminated, low, medium, and high concentration) required 80 ml of fresh goat milk. Precision was determined by calculating the relative standard deviation (RSD), while accuracy was assessed by calculating the % recovery from TPC testing of fresh goat milk using seven repetitions. The study was conducted from February to August 2024.

Fresh goat milk samples, each with a volume of 80 ml, were placed in four Erlenmeyer flasks (high, medium, low concentrations, and uncontaminated) containing 720 ml of BPW. The samples were then divided into eight Erlenmeyer flasks, each containing 90 ml for high concentration, eight Erlenmeyer flasks for medium concentration, eight for low concentration, and eight for uncontaminated samples. All samples were sterilized at 121°C for 15 minutes.

Bacterial concentrations in 10 ml of sterile distilled water (bacterial stock) were prepared by inoculating one loop of 18-24-hour-old agar culture into sterile distilled water. The turbidity was adjusted to the McFarland 0.5 standard (10⁸ CFU/ml). Bacterial suspension dilutions of 10³ - 10⁷ CFU/ml were prepared. For bacterial spiking into sample matrices, a 10⁷ CFU/ml bacterial suspension was added to high concentration samples, 10⁵ CFU/ml to medium concentration samples, 10³ CFU/ml to low concentration samples, and sterile distilled water to uncontaminated samples. All samples were then processed according to SNI 2897:2008 standards.

TPC testing based on SNI 2897:2008 involves several steps. First, 25 ml of fresh goat milk (uncontaminated, low, medium, and high concentrations) is measured aseptically and added to 225 ml of 0.1% sterile BPW, creating a 10⁻¹ dilution. Next, 1 ml of the 10⁻¹ dilution is transferred using a sterile pipette into 9 ml of BPW to obtain a 10⁻² dilution. Further dilutions (10⁻³, 10⁻⁴, 10⁻⁵, etc.) are prepared similarly as needed.

Then, 1 ml of each dilution suspension is placed into duplicate sterile Petri dishes. Subsequently, 15-20 ml of PCA, cooled to 45°C ± 1°C, is added to each Petri dish. To mix the sample solution and PCA medium thoroughly, the dishes are gently rotated forward and backward or in a figure-eight motion and left to solidify. Incubation is performed at 34°C to 36°C for 24-48 hours with inverted plates. Specifically for dairy products, incubation is conducted at 32°C ± 1°C for 24-48 hours with inverted plates.

RESULTS AND DISCUSSIONS

The verification process was divided into four treatment groups: samples without spike addition, samples with low, medium, and high concentrations, all tested in duplicate. Colonies growing on the media were observed and counted using a colony counter. The results of the observations and colony counts are presented in the table below.

Table 1. Results of Bacterial Colony Count in Goat Milk Samples

| No | Colony Count |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | Uncontaminated |  | Low Microbial Concentration (10⁻¹ Dilution) |  | Medium Microbial Concentration (10⁻² Dilution) |  | High Microbial Concentration (10⁻⁵ Dilution) |  |
|  | 1 | 2 | 1 | 2 | 1 | 2 | 1 | 2 |
| 1 | 0 | 0 | 7 | 7 | 153 | 183 | 106 | 116 |
| 2 | 0 | 0 | 3 | 3 | 165 | 169 | 128 | 118 |
| 3 | 0 | 0 | 9 | 8 | 164 | 203 | 115 | 102 |
| 4 | 0 | 0 | 7 | 6 | 148 | 209 | 131 | 125 |
| 5 | 0 | 0 | 5 | 6 | 150 | 174 | 111 | 126 |
| 6 | 0 | 0 | 4 | 3 | 162 | 159 | 125 | 128 |
| 7 | 0 | 0 | 5 | 5 | 158 | 180 | 116 | 121 |
| 8 | 0 | 0 | 6 | 6 | 105 | 95 | 105 | 115 |

In Table 1, it can be seen that the dairy goat milk samples with spike addition exhibit varying colony growth from repetition 1 to repetition 8. This variation occurs because the analyte used in the test is a living organism, whose growth can be influenced by factors such as nutrients, pH, oxygen, temperature, and incubation time, making it difficult to control. Nevertheless, the variation in colony numbers is considered relatively stable as it remains within the same order of magnitude, thus not significantly affecting the results. The required performance characteristics that must be calculated for the verification of quantitative microbiological methods in food and feed matrices are precision and accuracy.

A. Precision

Precision refers to the degree of agreement between individual test results when the method is repeatedly applied to multiple samplings of a homogenized sample. The accuracy of an analytical method is usually expressed as the relative standard deviation (CV or Coefficient of Variation) from a series of measurements.The recommended procedure for estimating inter-laboratory precision is the Relative Standard Deviation (RSD):

RSD =i=1i=nlogai-logbi/xi22p

Keterangan :

logai-logbixi= The relative difference

between the logarithm results of duplicates

i = 1,2,…n

p = number of duplicate determinations

Based on the RSD formula above, the RSD and CV values for dairy goat milk samples with low, medium, and high microbial concentrations are as follows:

Table 2

RSD and CV Values of Dairy Goat Milk Samples.

| No | Dairy Goat Milk Sample | RSD Value | CV Value |
| --- | --- | --- | --- |
| 1 | Low microbial concentration | 0,0259 | 2,59% |
| 2 | Medium microbial concentration | 0,0125 | 1,25% |
| 3 | High microbial concentration | 0,0037 | 0,37% |

Based on the table above, the RSD value for low microbial concentration is 0.0259, with a CV of 2.59%. For medium concentration, the RSD value is 0.0125, and the CV is 1.25%. For high concentration, the RSD value is 0.0037, and the CV is 0.37%. The acceptable RSD criterion used as a reference is < 5%. Therefore, the TPC testing method based on SNI 2897:2008 for dairy goat milk with low, medium, and high microbial concentrations can be

applied in the Integrated Laboratory of Poltekkes Kemenkes Denpasar.

B. Accuracy

Accuracy refers to the ability of a method to measure the true value or actual value of an analyte, such as the target microorganism. If the analyte is naturally present in the sample or intentionally added as part of a proficiency test, the method must be able to detect or recover the analyte at the correct concentration or frequency to be considered accurate. Accuracy is expressed as the percentage of analyte recovery from the added spike.

Table 3

Accuracy Test of Different Spike Colonies (SK)

Table 4

Accuracy (% Recovery) of TPC Testing in Goat Milk

| No | Dairy Goat Milk Sample | Average Log CFU | % Recovery Value |  |
| --- | --- | --- | --- | --- |
| 1 | Low microbial concentration (10⁻² CFU) | 1,72722 | 86,36114 |  |
| 2 | Medium microbial concentration (10⁻⁴ CFU) | 4,20308 | 105,0771 |  |
| 3 | High microbial concentration (10⁻⁷ CFU) | 7,07070 | 101,0101 |  |

Based on Tables 3 and 4, the percentage recovery values for TPC testing in dairy goat milk with low, medium, and high bacterial concentrations fall within the acceptable recovery percentage range (85–115%). Therefore, the TPC testing results using SNI 2897:2008 for dairy goat milk are considered accurate.

The verification of the TPC testing method in dairy goat milk is conducted to ensure that a method can be applied in the laboratory under real conditions through testing and proving that specific requirements have been met. According to Kartikasari (2021), accuracy and precision determination should be performed with seven repetitions. However, in this study, eight repetitions were conducted to improve accuracy and minimize systematic errors during the analysis stage.

Conclusion

The Total Plate Count (TPC) testing method according to SNI 2897:2008, using a dairy goat milk matrix, can be applied as a testing method in the Integrated Laboratory of Poltekkes Kemenkes Denpasar.

Recommendations

Future researchers can conduct verification of the Total Plate Count (TPC) testing method according to SNI 2897:2008 using different matrices.

References

- Faridah,DN.Erawan D, Sutriah K, Hadi A, Budiantari F. 2018. Implementasi SNI ISO/IEC 17025:2017 Persyaratan Umum Kompetensi Laboratorium Pengujian dan Laboratorium Kalibrasi. Jakarta : BSN

- Astuti DA, Sudarman A. 2012. Dairy Goats in Indonesia: Potential, Opportunities and Challenges. Proceedings of the 1st Asia Dairy Goat Conference

- Suguna, M., Rajeev Bhat, Wan Nadiah, W.A. 2012. Microbiological quality evaluation of goat milk collected from small scale dairy farms in Penang Island, Malaysia. Int. Food Res. J. 19 (3) 1241-1245

- Nurliyani, Suranindyah Y, Pretiwi P. 2015. Quality and Emulsion Stability of Milk from Ettawah Crossed Bred Goat During Frozen Storage. Procedia Food Sci. 3 142–149.

- Wardhani dan Nurbayanti . 2017 .VALIDASI METODE SNI 06-6989.12-2004 PADA PENETAPAN KESADAHAN TOTAL DALAM AIR PERMUKAAN SECARA KOMPLEKSIOMETRI .Buletin Teknik Litkayasa Akuakultur, 15 (2), 2017, 57-62

- Sanders, E. R. 2012. Aseptic Laboratory Techniques: Plating Methods. Journal of Visualized Experiments, (63), pp. 1–18. doi: 10.3791/3064.

- Singapore Accreditation Council. 2019. Guidance Notes C&B And ENV 002 Method Validation of Microbiological Methods